SAMPLE Protocol

PEDIATRIC POLYSOMNOGRAPHY PROTOCOL

Effective Date: 00/00/0000 (month/day/year)

Revision Date:

Approval Signature: To be signed by the site director

PURPOSE

Provide adequate information to perform quality polysomnograms on pediatric patients per physician order with confidence and competence. Pediatric patients include age-specific ranges of toddler (1–3 years), preschool child (4–5 years), school-age child (6–12 years), and adolescent (13–17 years). Polysomnography is essential to diagnose certain sleep-related conditions. Below is a list of common reasons:

- snoring

- excessive daytime somnolence

- insomnia, unexplained, resistant to therapy

- neuromuscular disease

- Disorders of arousal include, but are not limited to: confusional arousals, sleepwalking and sleep terrors. Other parasomnias include hypnagogic hallucinations, sleep paralysis, nocturnal seizures, bruxism, and rhythmic movement disorder.

- sleep-associated seizures

POLICY

Toddler (1–3 years), preschool (4–5 years), and school-age child (6–12 years)

- Schedule for nighttime testing with one parent or caretaker available throughout the procedure. A visit to the sleep facility prior to the scheduled procedure is encouraged to help parent and child become familiar with the surroundings and to answer any questions they might have.

- Encourage parent(s) to bring patient’s comfort items such as snacks, bottles, diapers and favorite bedtime toy or blanket.

- Patient-testing suite will be “child proofed” (outlet covers, no sharp corners or objects).

- Bed style will be determined based on size and age of patient. Cribs will be used for younger patients. Bedside safety rails will be used when appropriate. Protect mattress with the addition of waterproof pads to the normal bed linen in the event of risk of enuresis episodes.

- Parent(s) will stay in the patient suite at the discretion of the facility director or medical staff member. Parent(s) should not sleep in the same bed as the child, unless necessary to complete study. A recliner or bed should be provided for parents’ comfort, and whenever possible parents should be able to remain in the testing room.

- Feedings depend on the patient’s schedule. The fussy patient may do better if fed prior to electrode application. Feeding time may vary, depending on indications for and timing of the study.

- Additional toys and children’s videos will be available to keep patient occupied during electrode application.

- Due to some children’s apprehension to normal hospital uniforms, technologists may wear street clothes.

- Patient-to-technologist ratio will be 1:1, if indicated. The maximum ratio is 2:1.

- Electrodes will be applied to obtain consistency and maintain high-quality signals during the recording. Three methods of application are acceptable. The paste application method is preferred. The use of collodion is acceptable. For the patient with a history of reactive airways disease, history of allergy to collodion, an immediate reaction to collodion, or sensitivity to the EC-2 paste, the head-wrap method with Kling or Coban gauze may be used.

Adolescent (13–17 years)

- Schedule for nighttime testing with one parent or caretaker available throughout the procedure.

- Follow adult protocol for polysomnography testing.

- The maximum patient-to-technologist ratio is 2:1. In cases that involve significant anxiety or developmental disabilities, a 1:1 patient-to-technologist ratio may be necessary.

PROCEDURE

Prior to patient arrival at the sleep facility:

- Confirm physician order for polysomnogram and any other procedures such as supplemental nocturnal oxygen and/or nasal CPAP.

- Confirm that a history and physical examination are in the patient’s chart. Determine the child’s usual sleep/wake schedule and make an effort to schedule the polysomnogram start time within 30 minutes of the child’s usual bedtime.

- Confirm physician order or facility protocol for appropriate montages to be used.

- Confirm that age-specific resuscitation equipment is available in the sleep lab.

- Patient suite and sensors should be inspected and prepared prior to patient’s arrival.

- Calibrate polygraph and/or computer and related monitors and confirm proper operation.

- Confirm proper operation of video camera.

Patient/parent preparation for testing:

- After escorting the parent and the patient to the patient suite, explain procedure to the patient and parent in terms the child can understand.

- Consent-to-treat form should be signed.

Electrode placement/application:

- WASH HANDS THOROUGHLY

- EEG

  - Preparation of scalp electrodes:

    - The head must be measured according to the international 10-20 system to ensure accurate electrode application. Electrode sites generally used for the polysomnogram are ground, F3, F4, C3, C4, O1, O2, M1 and M2. Additional electrodes will depend on system or montages to be used. A system reference electrode is also used for computerized systems.

    - Prepare the electrode sites using skin abrasive specifically designated for skin preparation. Apply abrasive to marked electrode sites with a cotton swab using care to avoid overabrading the scalp. Allow it to dry.

    - Application of electrodes can be accomplished by a variety of techniques, either by use of TEN20 conductive paste and/or EC-2 cream or collodion glue. The uncooperative younger patient may require the head wrapping technique using Kling or Coban gauze. Using the paste application method is preferred.

    - Place TEN20 conductive paste in electrode cups and apply them to the prepared areas on the scalp. Secure the electrode cups by applying 1" × 2" gauze pieces with a small amount of EC-2 cream over each of the electrode cups and securing them with a few hairs crossed over the gauze patches.

- EOG

  - Properly identify each location of electrode placement, prepare sites with skin abrasive gel, and allow to dry.

  - Fill electrode cups with conductive cream and place the left electrooculogram electrode (EOG) 0.5-1 cm below and 0.5-1 cm lateral to the left outer canthus. Place the right electrooculogram electrode (EOG) 0.5-1cm above and 0.5-1 cm lateral to the right outer canthus.

  - Tape the electrodes in place. Guide wires toward the back of the patient’s neck.

- EMG

  - Submental/mentalis

    - Place three electrodes to record chin EMG: one in the midline 1 cm above the inferior edge of the mandible, one 1-2 cm below the inferior edge of the mandible and 1-2 cm to the right of the midline, one 1-2 cm below the inferior edge of the mandible and 1-2 cm to the left of the midline.

    - Secure electrodes with paper tape and drape over patient’s ears toward back of their neck.

  - Left/right anterior tibialis

    - Prepare each electrode site with skin abrasive gel. Allow it to dry.

    - Identify muscle by flexing patient’s ankle and observing muscle movement at the site.

    - Fill electrode cups with conductive cream and apply two electrodes 2-5 cm apart on that portion of patient’s legs. Secure with paper tape.

- EKG

  - Identify and prepare sites with skin abrasive gel and allow to dry.

  - Apply EKG electrodes or patches and attach snap electrodes. Placement of electrodes should be below the clavicles and equidistant from the sternum.

Respiration:

- Airflow devices

  - A single, double, or triple bead nasal and/or oral thermocouple or thermistor is secured directly below the nostrils to record air temperature changes at patient’s nose and/or mouth. A number of different devices are available. Nasal pressure transducers should be used, in accordance with the most current version of the AASM Scoring Manual. Choice of device is at the technologist’s discretion.

- CO2 monitoring (end-tidal or transcutaneous)

  - When evaluating sleep-disordered breathing in children, it is recommended that both airflow and CO2 be recorded.

- Respiratory effort

  - For monitoring respiratory effort, use one of the following:

  - Esophageal manometry

  - Dual thoracoabdominal RIP belts (calibrated or uncalibrated)

- Oximetry

  - An oximeter probe is attached to patient’s finger, toe, nose, or ear; choice of site depends on technologist discretion and available sensors.

Final preparation:

- Group sensor wires together and secure. Arrange wires out of reach of the child.

- Confirm patient’s comfort in bed.

- If applicable, raise the side rails to the up-and-locked position. Document in technologist’s notes.

- Confirm impedances.

- Confirm data at polygraph or computer.

Collection:

- Begin recording at this time and perform patient calibrations.

- After completing patient calibrations document “Lights Out.”

- Document all observed actions and reactions, important to include presence of snoring, movements or unusual and/or parasomnia behavior.

- Correct any artifact, and document it as it occurs during the recording.

- Sleep study data should generally be collected for a minimum of 8 hours, if indicated.

End of sleep study:

- At the end of the polysomnogram, gently awaken patient and document “Lights On.”

- End recording and if necessary exit from computer program.

- Gently remove all sensors from the patient. Take care to avoid irritation of patient’s sensitive skin.

- Carefully and diligently soak each electrode site with warm water until the electrode gently lifts away from the patient’s skin.

- Ensure that all paste residue has been removed by using a fine-toothed comb through the hair after all electrodes have been removed.

- When the parent (caretaker) and patient are ready to leave, escort family to sleep lab exit and document all observations and actions.

After the polysomnogram:

- Carefully sort wires and group them together by lengths and application sites.

- Remove any remaining tape, wash electrodes with soap and water, rinse and allow to soak in sodium hypochlorite solution for a minimum of 10 minutes. Rinse well and allow it to dry.

- Inspect wires at this time to ensure their integrity.

- Return any equipment and all cleaned and disinfected wires to their storage area for future use.

General cleanup checklist:

- Discard all used tape, collars, gauze, etc.

- Clean and disinfect thermocouples/thermistors, if not disposable.

- Return patient-preparation kit to appropriate area.

- Stock patient-preparation kit as needed.

- If CPAP and/or oxygen equipment was used, remove and empty humidifier, connecting tubing, nasal cannula, patient bore tubing and any other equipment and place in designated “dirty equipment area” for cleaning and disinfecting.

- Discard disposable equipment such as the nasal cannula or disposable oximeter probe.

- Remove any lint from CPAP equipment filter.

- Remove used linen and place in appropriate dirty linen container.

- Leave patient suites in clean and orderly condition.

Scoring the polysomnogram:

- Sleep-stage scoring is based on the most current version of the AASM Scoring Manual, if age appropriate.